# BeAST System

# Sensor Processing & Outcome Metrics Summary

## Overview

This document summarizes all sensor data processing pipelines implemented in the BeAST system, from raw sensor output through final database variables. All outcome metrics are normalized to 0-100 scales for consistency and understanding for command/warfighters.

## 📊 QUICK REFERENCE: UNIVERSAL SCALES

All metrics normalized to 0-100 scale with 4-zone classification:

- Zone 1 (Green): Optimal performance - No intervention

- Zone 2 (Yellow): Mild risk - Monitor closely

- Zone 3 (Orange): Moderate risk - Consider intervention

- Zone 4 (Red): High risk - Immediate action required

Exception: Some metrics are "lower is better" (fatigue, stress) while others are "higher is better" (attention, hydration). Zone definitions account for this directionality.

## 📊 SENSOR 1: EEG (12 Channels + 2 Reference + 2 Bias)

### Raw Data Input

| Parameter | Unit | Range | Sample Rate |
| --- | --- | --- | --- |
| Raw EEG Voltage | mV (millivolts) | ±0.1 mV (±100 µV) | 250-500 Hz |
| Channel Count | - | 12 active + 2 ref + 2 bias | - |
| ADC Resolution | bits | 16-24 bit | - |

### Data Processing Steps (Simplified)

- Acquisition

  - Collect raw mV data from 12 channels

  - Use 2 reference channels for differential measurement

  - BIAS channels for noise cancellation (removed from analysis)

- Preprocessing

  - DC offset removal (subtract mean)

  - Notch filter (50/60 Hz power line noise)

  - Bandpass filter (0.5-50 Hz)

  - Re-referencing (Common Average Reference)

- Artifact Removal

  - Motion artifacts (IMU-based adaptive filtering)

  - Eye blinks (ICA or template matching)

  - Muscle artifacts (>10 µV threshold rejection)

  - Electrode quality monitoring

- Frequency Band Extraction (Per Channel)

  - Fast Fourier Transform (FFT) with Welch's method

  - Extract power spectral density for each band:

    - Delta: 0.5-4 Hz

    - Theta: 4-8 Hz

    - Alpha: 8-13 Hz

    - Beta: 13-30 Hz

    - Gamma: 30-45 Hz

- Feature Calculation

  - Band power per channel (µV²)

  - Aggregate across channels (mean/median)

  - Calculate band ratios

  - Compute engagement indices

### Outcome Variables (0-100 Scale)

| Variable Name | Original Unit/Scale | Conversion to 0-100 | Clinical Range | Zone Definitions |
| --- | --- | --- | --- | --- |
| Alpha/Theta Ratio - Fatigue | Ratio (0.5-2.5) | 100 × ((ATR - 0.5) / 2.0) | Optimal: >1.5, Fatigue: <0.7 | Z1: 80-100, Z2: 60-79, Z3: 40-59, Z4: 0-39 |
| Beta/Alpha Ratio | Ratio (0.5-3.0) | 100 × ((BAR - 0.5) / 2.5) | Normal: 0.8-1.5, Overload: >2.0 | Z1: 20-60, Z2: 61-75, Z3: 76-85, Z4: >85 |
| Theta/Beta Ratio | Ratio (0.3-3.0) | 100 × (1 - ((TBR - 0.3) / 2.7)) | Alert: <0.5, Drowsy: >2.5 | Z1: 80-100, Z2: 60-79, Z3: 40-59, Z4: 0-39 |
| Engagement Index | Ratio (0.1-2.0) | 100 × ((EI - 0.1) / 1.9) | Optimal: >1.0, Poor: <0.2 | Z1: 75-100, Z2: 50-74, Z3: 25-49, Z4: 0-24 |
| Cognitive Load | Composite (0-1) | 100 × CL | Low: <0.3, High: >0.7 | Z1: 0-30, Z2: 31-50, Z3: 51-70, Z4: 71-100 |
| Attention Focus | Composite (0-1) | 100 × AF | Optimal: >0.7, Poor: <0.3 | Z1: 70-100, Z2: 50-69, Z3: 30-49, Z4: 0-29 |
| Attention Capacity | Percentage | Direct (0-100%) | Optimal: >80%, Critical: <30% | Z1: 80-100, Z2: 60-79, Z3: 40-59, Z4: 0-39 |
| Microsleep Detection | Binary/Frequency | Events per 15-minute window | Normal: 0, Warning: 1-2, Critical: >3 | Z1: 0, Z2: 1, Z3: 2, Z4: ≥3 |

### Database Schema Variables

Table: eeg_band_power

- delta_power_uv2 (per channel, µV²)

- theta_power_uv2 (per channel, µV²)

- alpha_power_uv2 (per channel, µV²)

- beta_power_uv2 (per channel, µV²)

- gamma_power_uv2 (per channel, µV²)

Table: eeg_calculated_metrics

- alpha_theta_ratio (dimensionless)

- beta_alpha_ratio (dimensionless)

- theta_beta_ratio (dimensionless)

- engagement_index (dimensionless)

- cognitive_load_score (0-100)

- attention_focus_score (0-100)

- attention_capacity_score (0-100)

## 📊 SENSOR 2: PPG (Photoplethysmography - MAXM86161)

### Raw Data Input

| Parameter | Unit | Range | Sample Rate |
| --- | --- | --- | --- |
| Red LED counts | ADC counts | 0-524,287 (19-bit) | 25-400 Hz |
| Infrared LED counts | ADC counts | 0-524,287 (19-bit) | 25-400 Hz |
| Green LED counts | ADC counts | 0-524,287 (19-bit) | 25-400 Hz |
| Ambient Light Sensor | ADC counts | 0-65,535 (16-bit) | 1-10 Hz |
| Die Temperature | °C | -40 to +85°C | 1 Hz |

### Data Processing Steps (Simplified)

- ADC to Physical Units Conversion

  - Convert ADC counts to µA or absorption units

  - Temperature compensation using die temp

  - Ambient light correction

- Signal Preprocessing

  - Bandpass filter: 0.5-5 Hz (cardiac frequency range)

  - Moving average smoothing (3-5 sample window)

  - Baseline drift removal

- Peak Detection

  - Adaptive threshold algorithm

  - Systolic peak identification

  - Inter-Beat Interval (IBI) extraction

  - Ectopic beat rejection (>20% deviation)

- SpO2 Calculation

  - AC/DC ratio for Red and IR channels

  - Ratio of ratios: R = (AC_red/DC_red) / (AC_ir/DC_ir)

  - Empirical calibration: SpO2 = 110 - 25×R

- Heart Rate Variability (HRV)

  - Time-domain metrics (RMSSD, SDNN, pNN50)

  - Frequency-domain (FFT): LF (0.04-0.15 Hz), HF (0.15-0.4 Hz)

  - Non-linear metrics (Poincaré, entropy)

- Derived Metrics

  - Stress Index: SI = AMo / (2 × MxDMn × Mo)

  - Perfusion Index: AC/DC ratio

  - Cardiac Output estimation

  - Autonomic balance (LF/HF ratio)

### Outcome Variables (0-100 Scale)

| Variable Name | Original Unit/Scale | Conversion to 0-100 | Clinical Range | Zone Definitions |
| --- | --- | --- | --- | --- |
| Heart Rate | BPM | See zones | Rest: 60-100, Max: 220-age | Z1: 60-100, Z2: 101-120, Z3: 121-150, Z4: >150 or <50 |
| SpO2 | Percentage | Direct (85-100%) | Normal: 95-100%, Critical: <90% | Z1: 95-100, Z2: 90-94, Z3: 85-89, Z4: <85 |
| HRV RMSSD | ms | 100 × (RMSSD / 100) | Healthy: 20-100ms, Poor: <10ms | Z1: 50-100, Z2: 30-49, Z3: 15-29, Z4: 0-14 |
| Stress Index | Arbitrary units | 100 × (SI / 500) capped at 100 | Normal: 50-150, High: >300 | Z1: 0-30, Z2: 31-60, Z3: 61-80, Z4: 81-100 |
| Physical Fatigue | Composite | 100 × (1 - [RMSSD_current / RMSSD_baseline]) | Based on HRV degradation | Z1: 0-25, Z2: 26-50, Z3: 51-75, Z4: 76-100 |
| Perfusion Index | Percentage | 100 × (PI / 20) capped at 100 | Normal: 2-20%, Poor: <1% | Z1: 40-100, Z2: 20-39, Z3: 10-19, Z4: 0-9 |
| Autonomic Balance (LF/HF) | Ratio | 100 × (LF_HF / 5) capped at 100 | Balanced: 0.5-2.0, Stress: >3.0 | Z1: 10-40, Z2: 41-60, Z3: 61-80, Z4: >80 |

### Database Schema Variables

Table: ppg_raw_data

- red_counts (ADC, integer)

- ir_counts (ADC, integer)

- green_counts (ADC, integer)

- als_counts (ADC, integer)

- die_temp_c (°C, float)

Table: ppg_calculated_metrics

- heart_rate_bpm (integer)

- spo2_percent (float, %)

- perfusion_index (float, %)

- signal_quality (float, 0-1)

Table: hrv_metrics

- rmssd_ms (float, ms)

- sdnn_ms (float, ms)

- pnn50_percent (float, %)

- lf_power_ms2 (float, ms²)

- hf_power_ms2 (float, ms²)

- lf_hf_ratio (float, dimensionless)

- stress_index (float, arbitrary units)

## 📊 SENSOR 3: Bioimpedance (12 Channels + 2 Reference + 2 Bias)

### Raw Data Input

| Parameter | Unit | Range | Sample Rate |
| --- | --- | --- | --- |
| In-phase (I) component | ADC counts | Varies by chip | 50-100 Hz |
| Quadrature (Q) component | ADC counts | Varies by chip | 50-100 Hz |
| Excitation Frequency | kHz | Typically 50 kHz | Fixed |
| Reference Amplitude | mA or V | Chip-specific | Fixed |

### Data Processing Steps (Simplified)

- Demodulation

  - Extract I and Q components from ADC

  - Apply reference calibration factors

- Complex Impedance Calculation (Per Channel)

  - Magnitude: Z = sqrt(I² + Q²) [Ohms]

  - Phase Angle: φ = arctan(Q/I) [degrees]

  - Resistance: R = Z × cos(φ) [Ohms]

  - Reactance: X = Z × sin(φ) [Ohms]

- Regional Aggregation

  - Group channels by anatomical region:

    - Left ear: Channels 1-3

    - Right ear: Channels 4-6

    - Cross-head: Channels 7-9

    - Auxiliary: Channels 10-12

- Feature Extraction

  - Mean Phase Angle: Average across channels/regions

  - Impedance Magnitude: Mean resistance

  - Phase Angle Dispersion: Variation across channels

  - Temporal Changes: Rate of change over time

- Clinical Metric Derivation

  - Hydration Status: Based on phase angle thresholds

  - Total Body Water (TBW): Empirical equations using R and height/weight

  - Extracellular/Intracellular Water Ratio

  - Dehydration Risk Score: Multi-factor composite

### Outcome Variables (0-100 Scale)

| Variable Name | Original Unit/Scale | Conversion to 0-100 | Clinical Range | Zone Definitions |
| --- | --- | --- | --- | --- |
| Phase Angle Mean | Degrees (3-10°) | 100 × ((PA - 3) / 7) | Healthy: >5°, Poor: <4° | Z1: 71-100 (>6°), Z2: 57-70 (5-6°), Z3: 43-56 (4-5°), Z4: 0-42 (<4°) |
| Hydration Status | Percentage | Based on TBW% of body weight | Normal: 55-65%, Dehydrated: <50% | Z1: 78-100 (55-65%), Z2: 71-77 (50-54%), Z3: 64-70 (45-49%), Z4: 0-63 (<45%) |
| Dehydration Risk Score | Composite (0-10) | 100 × (DRS / 10) | Low: 0-3, High: 7-10 | Z1: 0-30, Z2: 31-50, Z3: 51-70, Z4: 71-100 |
| Perfusion Index (Bio) | Percentage | 100 × (PI / 20) | Normal: 5-15%, Poor: <3% | Z1: 50-100, Z2: 30-49, Z3: 15-29, Z4: 0-14 |
| Metabolic Stress Index | Composite | Complex formula from R, PA, HR | Normal: <40, High: >70 | Z1: 0-40, Z2: 41-60, Z3: 61-80, Z4: 81-100 |

### Database Schema Variables

Table: bioimpedance_raw_data

- channel_id (1-12, integer)

- impedance_magnitude_ohms (float, Ω)

- phase_angle_degrees (float, °)

- resistance_ohms (float, Ω)

- reactance_ohms (float, Ω)

- signal_quality_index (float, 0-1)

Table: bioimpedance_clinical_metrics

- phase_angle_mean (float, °)

- hydration_status_percent (float, %)

- total_body_water_percent (float, %)

- dehydration_risk_score (float, 0-100)

- perfusion_index (float, %)

- metabolic_stress_index (float, 0-100)

## 📊 SENSOR 4: IMU (Inertial Measurement Unit - 2 Devices)

### Raw Data Input

| Parameter | Unit | Range | Sample Rate |
| --- | --- | --- | --- |
| Accelerometer (3-axis) - Normal Mode | g or m/s² | ±16g typical | 100-1000 Hz |
| Accelerometer (3-axis) - High-G Mode | g or m/s² | >16g (extended range) | 100-1000 Hz |
| Gyroscope (3-axis) | °/s or rad/s | ±2000°/s typical | 100-1000 Hz |
| Magnetometer (3-axis) | µT or Gauss | Earth field ±50µT | 50-200 Hz |

### CRITICAL FEATURE: Dual-Range Accelerometer

### Normal Mode (±16g): Continuous monitoring for movement, gait, posture

### High-G Detection (>16g): Separate detection path triggered for blast/impact events

### Captures extreme acceleration events (blasts, crashes, falls from height)

### Prevents sensor saturation during high-energy impacts

### Critical for TBI screening and cumulative blast exposure tracking

### Often uses dedicated interrupt or secondary data path in IMU hardware

### Data Processing Steps (Simplified)

- Calibration

  - Zero-offset correction

  - Sensitivity adjustment

  - Temperature compensation

  - Cross-calibration between normal and high-G modes

- Dual-Path Data Handling

- Normal Path (<±16g): Continuous streaming for routine monitoring

- High-G Path (>16g): Event-triggered capture with:

  - Pre-trigger buffering (50-100ms before event)

  - Extended capture duration (500ms-1s post-event)

  - High-resolution timestamp

  - Peak value detection and hold

- Sensor Fusion

  - Quaternion-based orientation estimation

  - Madgwick or Mahony algorithms

  - Combine accel, gyro, mag into 3D orientation

- Feature Extraction

  - Normal Mode Features:

    - Postural Sway: Standard deviation of position

    - Gait Analysis: Step regularity, cadence

    - Head Movement Patterns: Speed, frequency, smoothness

    - Fall Detection: Rapid orientation changes

  - High-G Event Features:

- Peak acceleration magnitude (max g-force)

- Impact direction vector

- Jerk (rate of acceleration change: dA/dt)

- Energy absorption (integral of acceleration over time)

- Duration above threshold (time >16g, >50g, >100g)

- Impact Classification

  - Blast Events: Characteristic pressure wave signature

  - Vehicle Crashes: Sustained multi-axis acceleration

  - Falls: Vertical component dominance + rotation

  - Weapon Recoil: Directional pattern recognition

  - False Positives: Filter out equipment drops, door slams

- TBI Correlation/Risk Assessment

  - Cross-reference high-G events with immediate EEG changes

  - Compare impact characteristics to known TBI thresholds:

    - Linear acceleration: >80-100g for >2ms = possible concussion

    - Rotational acceleration: >4500 rad/s² = high TBI risk

  - Cumulative exposure dosimetry

  - Return-to-duty decision support

- Cumulative Exposure Tracking

- Log all events >16g with timestamp and magnitude

- Calculate daily/weekly/deployment exposure totals

- Track frequency and intensity trends

- Flag concerning patterns (multiple sub-concussive events)

### Outcome Variables (0-100 Scale)

| Variable Name | Original Unit/Scale | Conversion to 0-100 | Clinical Range | Zone Definitions |
| --- | --- | --- | --- | --- |
| Peak Acceleration (Normal) | g-forces | 100 * (accel / 16) | Movement: 0-2g, Activity: 2-8g, Impact: >8g | Z1: 0-12.5 (0-2g), Z2: 12.6-50 (2-8g), Z3: 51-75 (8-12g), Z4: 76-100 (12-16g) |
| Peak Acceleration (High-G) | G-forces (>16g) | See Zones Below | TBI risk: >80gg, Severe: >100g | Z1: 0-10g, Z2: 11-50g, Z3: 51-100g, Z4: >100g |
| High-G Event Count | Events per period | Direct count | Safe: 0-1/day, Concern: >2/day | Z1: 0, Z2: 1-2/day, Z3: 3-5/day, Z4: >5/day |
| Cumulative Blast Exposure | Psi-ms (overpressure * duration) | Sum over period | Low: <20, High: >100 | Z1: 0-20, Z2: 21-50, Z3: 51-100, Z4: >100 |
| Movement Quality Index | Composite (0-1) | 100 × MQI | Optimal: >0.8, Poor: <0.4 | Z1: 80-100, Z2: 60-79, Z3: 40-59, Z4: 0-39 |
| Postural Stability | Composite (0-1) | 100 × (1 - Sway_Score) | Stable: >0.7, Unstable: <0.3 | Z1: 70-100, Z2: 50-69, Z3: 30-49, Z4: 0-29 |
| Gait Irregularity | Coefficient of Variation | 100 × CV | Normal: <10%, Abnormal: >20% | Z1: 0-10, Z2: 11-15, Z3: 16-20, Z4: >20 |

### Database Schema Variables

Table: imu_raw_data

- sensor_location (text: left_ear, right_ear, helmet)

- accel_x, accel_y, accel_z (float, g or m/s²)

- gyro_x, gyro_y, gyro_z (float, °/s)

- mag_x, mag_y, mag_z (float, µT)

- impact_magnitude (float, g)

- blast_flag (boolean)

Table: imu_calculated_metrics

- peak_acceleration_g (float)

- cumulative_exposure_psi_ms (float)

- movement_quality_index (float, 0-100)

- postural_stability_score (float, 0-100)

- gait_irregularity_percent (float, %)

## 📊 SENSOR 5: Core Temperature

### Raw Data Input

| Parameter | Unit | Range | Sample Rate |
| --- | --- | --- | --- |
| Core Body Temperature | °C | 35-42°C typical | 1-10 Hz |
| Sensor Location | - | Tympanic/ear canal | - |

### Data Processing Steps (Simplified)

- Raw Temperature Reading

  - ADC conversion to °C

  - Sensor-specific calibration

- Smoothing

  - Moving average (10-30 second window)

  - Outlier rejection

- Rate of Change Analysis

  - dT/dt calculation

  - Trend detection (rising/falling)

- Circadian Modeling

  - Fit to expected diurnal rhythm

  - Detect abnormal deviations

- Heat Stress Correlation

  - Combine with ambient temp & humidity

  - Integrate with HR and sweating proxies

### Outcome Variables (0-100 Scale)

| Variable Name | Original Unit/Scale | Conversion to 0-100 | Clinical Range | Zone Definitions |
| --- | --- | --- | --- | --- |
| Core Temperature | °C | 100 × ((T - 36.5) / 3.5) | Normal: 36.5-37.5°C, Critical: >40°C | Z1: 0-28 (36.5-37.5°C), Z2: 29-57 (37.6-38.5°C), Z3: 58-85 (38.6-39.5°C), Z4: 86-100 (>39.5°C) |
| Heat Stress Index | Composite | Multi-factor: Temp + HR + environment | Normal: <30, Critical: >70 | Z1: 0-30, Z2: 31-50, Z3: 51-70, Z4: 71-100 |
| Temperature Rate of Change | °C/hour | Normalized to risk score | Normal: <0.5°C/hr, Rapid: >2°C/hr | Z1: 0-25, Z2: 26-50, Z3: 51-75, Z4: 76-100 |

### Database Schema Variables

Table: core_temperature

- temperature_celsius (float, °C)

- sensor_location (text)

- measurement_quality (float, 0-1)

Table: heat_stress_metrics

- core_temp_celsius (float)

- heat_stress_index (float, 0-100)

- temp_rate_of_change_c_per_hour (float)

## 📊 SENSOR 6: Environmental Monitoring

### Raw Data Input

| Parameter | Unit | Range | Sample Rate |
| --- | --- | --- | --- |
| Ambient Temperature | °C | -40 to +85°C | 0.1-1 Hz |
| Relative Humidity | % | 0-100% | 0.1-1 Hz |
| Sound Level | dB(A) | 30-140 dB | 1-10 Hz |
| Overpressure | psi or kPa | 0-50 psi | 100,000 Hz (during blast) |

### Data Processing Steps (Simplified)

- Temperature & Humidity

  - ADC to physical units

  - Calculate heat index: HI = function(T, RH)

  - WBGT estimation

- Acoustic Monitoring

  - A-weighted sound pressure level

  - Time-weighted average (TWA)

  - Peak exposure detection

  - Cumulative dose calculation

- Overpressure Detection

  - High-speed sampling during blast events

  - Peak pressure measurement

  - Impulse duration calculation

  - Integration: ∫ P(t) dt

- CBRN Integration (if equipped)

  - Agent detection thresholds

  - Exposure duration tracking

  - Correlation with physiological responses

### Outcome Variables (0-100 Scale)

| Variable Name | Original Unit/Scale | Conversion to 0-100 | Clinical Range | Zone Definitions |
| --- | --- | --- | --- | --- |
| Heat Index | °C or °F | Based on risk categories | Safe: <32°C, Danger: >41°C | Z1: 0-25 (<32°C), Z2: 26-50 (32-38°C), Z3: 51-75 (39-41°C), Z4: 76-100 (>41°C) |
| Acoustic Exposure Dose | % of limit | Direct percentage | Safe: <50%, Dangerous: >100% | Z1: 0-50, Z2: 51-75, Z3: 76-100, Z4: >100 |
| Blast Overpressure | psi | Based on TBI risk thresholds | Safe: <0.5, TBI risk: >4 | Z1: 0-12.5 (<0.5 psi), Z2: 12.6-25 (0.5-2 psi), Z3: 26-75 (2-4 psi), Z4: 76-100 (>4 psi) |

### Database Schema Variables

Table: environmental_data

- ambient_temp_c (float, °C)

- humidity_percent (float, %)

- heat_index (float, °C)

- sound_level_db (float, dB(A))

Table: overpressure_events

- peak_pressure_psi (float)

- impulse_psi_ms (float)

- duration_ms (float)

- blast_number (integer)

## 🔄 MULTI-MODAL DATA FUSION METRICS

### These metrics require integration of multiple sensors

| Outcome Variable | Input Sensors | Original Scale | 0-100 Conversion | Zone Definitions |
| --- | --- | --- | --- | --- |
| Composite Fatigue Index (CFI) | EEG (40%) + PPG (30%) + Temp (20%) + Bio (10%) | Weighted composite | Direct output | Z1: 0-20 (optimal), Z2: 21-40, Z3: 41-60, Z4: 61-100 (severe) |
| Cognitive Load | EEG (Theta/Beta) + PPG (HRV) | Weighted composite | Direct output | Z1: 0-30 (low), Z2: 31-50, Z3: 51-70, Z4: 71-100 (overload) |
| Tiredness | EEG (Alpha/Theta) + Temp + IMU (movement) | Weighted composite | Direct output | Z1: 0-20 (alert), Z2: 21-40, Z3: 41-60, Z4: 61-100 (drowsy) |
| Attention Focus | EEG (Engagement Index) + IMU (head movement) | Weighted composite | Direct output | Z1: 70-100 (optimal), Z2: 50-69, Z3: 30-49, Z4: 0-29 |
| Stress Index | PPG (HRV LF/HF) + EEG (Beta asymmetry) + Bio | Weighted composite | Direct output | Z1: 0-30 (low), Z2: 31-50, Z3: 51-70, Z4: 71-100 (acute) |
| Neurovascular Coupling | EEG (neural activity) + PPG (blood flow) | Correlation coefficient | 100 × (1 - abs(corr)) | Z1: 80-100 (tight coupling), Z2: 60-79, Z3: 40-59, Z4: 0-39 (uncoupled) |
| Metabolic Stress Index | PPG (HR, cardiac output) + Bio (hydration) + Temp | Weighted composite | Direct output | Z1: 0-40 (low), Z2: 41-60, Z3: 61-80, Z4: 81-100 (high) |
| Compensation Cognitive Load | EEG (increased frontal activity) + PPG (elevated HR) | Effort - Performance delta | Direct output | Z1: 0-30 (efficient), Z2: 31-50, Z3: 51-70, Z4: 71-100 (compensating) |
| Fatigue Severity Score | Multi-modal: All sensors + self-report correlation | Clinical FSS (1-7) | 100 × ((FSS - 1) / 6) | Z1: 0-33 (FSS 1-3), Z2: 34-50 (FSS 3-4), Z3: 51-66 (FSS 4-5), Z4: 67-100 (FSS >5) |
| Attention Capacity | EEG + PPG + task performance proxy | Percentage of maximum | Direct (0-100%) | Z1: 80-100, Z2: 60-79, Z3: 40-59, Z4: 0-39 |

## 📋 ON-GOING DEVELOPMENT - MISSING, NEEDING CONFIRMED OR UNDEFINED ELEMENTS

### Critical Gaps Identified

- Sampling Rates - (Need Confirmation at Sensor/in-ear level)

  - ✅ EEG: Specified (250-500 Hz)

  - ✅ PPG: Specified (25-400 Hz, typically 100 Hz)

  - ✅ Bioimpedance: Specified (50-100 Hz)

  - ✅ IMU: Specified (100-500 Hz)

  - ✅ Temperature: Specified (1-10 Hz)

  - ✅ Environmental: Specified (0.1-10 Hz depending on parameter)

- Electrode/Sensor Locations

  - ✅ EEG: 12 channels in ear canal and post-auricular, V7 locations

  - ✅ PPG: Left and right ear (6 channels, 1 BIAS, 1 REF per ear)

  - ✅ Bioimpedance: Same locations as EEG

  - ✅ IMU: 2 devices (near ears/helmet)

  - ⚠️ Missing: Exact channel mapping to anatomical positions with reference to Skull Cap comparisons.

- Assignment

- Session Start/Stop

- Calibration Parameters

  - ⚠️ Missing: Individual baseline establishment protocol (duration, conditions)

  - ⚠️ Missing: Re-calibration frequency and triggers

  - ⚠️ Missing: Calibration validation thresholds

  - ⚠️ Missing: Baseline data to device

- State-Based Adjustments (Not for December, will use defaults)

  - ⚠️ Missing: How sampling rates change with state transitions

  - ⚠️ Missing: Processing algorithm changes per state (normal, high-stress, sleep)

  - ⚠️ Missing: Alert threshold adjustments by state

  - ⚠️ Missing: Comparisons to warfighters baselines

- Data Quality Metrics

  - ✅ Signal quality indices defined for each sensor

  - ⚠️ Missing: Minimum data quality thresholds for metric calculation

  - ⚠️ Missing: Handling of missing/corrupted data segments

- Temporal Windows

  - ⚠️ Missing: Window sizes for moving averages on each metric

  - ⚠️ Missing: Minimum data duration required for reliable metric calculation

  - ⚠️ Missing: Look-back periods for trend analysis

- Inter-Sensor Synchronization (Sensor acquisition level)

  - ⚠️ Missing: Clock synchronization protocol

  - ⚠️ Missing: Maximum acceptable time skew between sensors

  - ⚠️ Missing: Resampling strategy for fusion (common time base)

- Zone Transition Logic

  - ⚠️ Missing: Hysteresis rules (preventing rapid zone oscillation)

  - ⚠️ Missing: Minimum duration in zone before recording transition

  - ⚠️ Missing: Multi-metric conflict resolution (when sensors disagree)

- Cumulative Risk Scoring

  - ⚠️ Missing: Time-weighted risk integration formula

  - ⚠️ Missing: Recovery credit calculation (how zones improve over rest)

  - ⚠️ Missing: Long-term trend weighting vs. acute events

- Performance Degradation Models

  - ⚠️ Missing: Task-specific performance prediction from sensor fusion

  - ⚠️ Missing: Individual vs. population-based normative curves

  - ⚠️ Missing: Mission-type specific threshold adjustments

- Connection Bridge/APIs for data transfer

  - ⚠️ Missing: API/Bridge for LIVE connection and data to dashboard

  - ⚠️ Missing: API/Bridge for data transfer of sessions from puck Db to external Db

  - ⚠️ Missing:

## 🔧 IMPLEMENTATION RECOMMENDATIONS

### For Programming Team

- Priority 1: Data Pipeline Architecture

  - Implement sensor-specific preprocessing classes

  - Create unified timestamp/synchronization system

  - Build quality control and artifact rejection pipeline

  - Establish baseline calculation and storage system

- Priority 2: Feature Extraction

  - Develop FFT/spectral analysis module for EEG

  - Build peak detection and HRV calculation for PPG

  - Implement complex impedance calculation for bioimpedance

  - Create sensor fusion engine for multi-modal metrics

- Priority 3: Scoring & Classification

  - Code 0-100 normalization functions for all metrics

  - Implement zone classification logic with hysteresis

  - Build alert generation system with prioritization

  - Create zone transition tracking and logging

- Priority 4: Database Integration

  - Map processed metrics to database schema

  - Implement efficient time-series storage

  - Create materialized views for real-time queries

  - Build data export functions for dashboard

- Priority 5: Validation & Testing

  - Unit tests for each processing function

  - Integration tests for sensor fusion

  - Validation against ground truth datasets

  - Performance benchmarking (latency, throughput)

### Next Steps Questions for Stakeholders

- Baseline Establishment: How long should initial baseline recording be? Under what conditions?

- Individual vs. Population Norms: What weight to give personal baseline vs. military population norms?

- Alert Fatigue: How to minimize false positives while maintaining safety? What confirmation logic?

- Data Retention: How long to keep raw data vs. processed metrics? Privacy considerations?

- Real-Time vs. Retrospective: Which metrics need <1s latency vs. which can use longer windows for accuracy?

- Edge Computing: What processing happens on-device vs. in cloud? Bandwidth constraints?

- Interoperability: Integration requirements with existing military medical systems?

## 📚 SUMMARY: COMPLETE VARIABLE LIST

### Final Database Outcome Variables (All 0-100 Scale)

| Category | Variable Name | Primary Sensor(s) | Fusion Required |
| --- | --- | --- | --- |
| Cognitive | Cognitive Load | EEG, PPG | Yes |
| Cognitive | Attention Focus | EEG, IMU | Yes |
| Cognitive | Attention Capacity | EEG, PPG | Yes |
| Cognitive | Compensation Cognitive Load | EEG, PPG | Yes |
| Fatigue/Alertness | Composite Fatigue Index (CFI) | EEG, PPG, Temp, Bio | Yes |
| Fatigue/Alertness | Tiredness | EEG, Temp, IMU | Yes |
| Fatigue/Alertness | Fatigue Severity Score | Multi-modal | Yes |
| Stress | Stress Index | PPG, EEG, Bio | Yes |
| Metabolic | Metabolic Stress Index | PPG, Bio, Temp | Yes |
| Neurovascular | Neurovascular Coupling Index | EEG, PPG | Yes |
| Cardiovascular | Heart Rate | PPG | No |
| Cardiovascular | SpO2 | PPG | No |
| Cardiovascular | HRV RMSSD | PPG | No |
| Cardiovascular | Physical Fatigue | PPG | No |
| Hydration | Hydration Status | Bioimpedance | No |
| Hydration | Dehydration Risk Score | Bioimpedance, PPG, Temp | Yes |
| Hydration | Phase Angle Mean | Bioimpedance | No |
| Thermal | Core Temperature | Temperature | No |
| Thermal | Heat Stress Index | Temperature, Environment | Yes |
| Motion | Movement Quality Index | IMU | No |
| Motion | Postural Stability | IMU | No |
| Environmental | Heat Index | Environment | No |
| Environmental | Acoustic Exposure Dose | Environment | No |
| Environmental | Blast Overpressure | Environment, IMU | Yes |

Total Outcome Variables: 24
 Requiring Multi-Modal Fusion: 14
 Single-Sensor Derived: 10